## Sheet1
Signal and signal-to-noise ratio of absorption spectrophotometry

Input Data			Calculated Outputs
Light source
Blackbody temperature (Tc)	2227	C	Blackbody temperature (Tk)	2500	K
Emissivity (emisiv)	0.5		Total radiance (B)	35.25	W/sr/cm2
Flicker noise fraction (flicker)	0.00%		Spectral radiance at exw (Blambda)	3.30E-04	W/sr/nm/cm2
			Peak wavelength	1158.8	nm
			Energy of photon at exw (e)	4.97E-19	Joules

Monochromator
Wavelength setting (exw)	400	nm	Slit area (SlitAex)	0.050	cm2
Slit height (Hex)	5	mm	Radiant power input (PhiInEx)	4.23E-02	watts
Slit width (Wex)	1	mm	Spectral bandpass (SBex)	2.80	nm
Transmission factor of monochromator optics (Tex)	0.5		Radiant power output (PhiOutEx)	5.55E-07	watts
Solid angle (Sax)	0.024	sr
Reciprocal linear dispersion (RLDex)	2.8	nm/mm
Fractional stray light (sl)	0.00%

Analytical Sample				Sample cell	Reference cell
Molar absorptivity of analyte (a)	1.000	liters/mole-cm	Absorbance at measured wavelength (abs, absRef)	0.0010	0.0000
Path length of cell (l)	1.000	cm	Fraction absorbed (F, Fref)	0.230%	0.000%
Molar concentration of analyte ( c )	0.001	moles/liter	Radiant power transmitted through sample (PhiAbs, PhiRef)	5.54E-07	5.55E-07	watts
Solvent absorbance at exw (sa)	0

Detector (photomultiplier tube)				Sample cell	Reference cell
			Cathode sensitivity (Rlam)	0.064	0.064	amp/watt
Secondary emission gain per stage (g)	3.21		Radiant power on photocathode (PhiDet, PhiRef)	5.54E-07	5.55E-07	watts
Quantum efficiency of cathode (Klam)	0.2		photocathodic current (Ic)	3.57E-08	3.57E-08	amps
photomultiplier gain (m)	37000		photoanodic current (Ie, Ier)	1.32E-03	1.32E-03	amps
cathode thermionic emission rate (Rt)	1.00E+04	Sec-1	Photon noise rms current (PhotNoise, PhotNoiseRef)	3.37E-09	3.37E-09	amps
Noise bandwidth (deltaf)	0.5	Sec-1	Themionic anode current (Iat)	5.93E-11	5.93E-11	amps
			Thermionic noise rms current (ThermNoise)	7.14E-13	7.14E-13	amps
			Flicker noise current (FlickerNoise, FlickerNoiseRef)	0.00E+00	0.00E+00	amps
			Total rms noise (TotNoise, TotNoiseRef)	3.37E-09	3.37E-09	amps
			Intensity signal-to-noise ratio (SNR)	391483.7	391934.7
			Percent photon noise	99.98%	99.98%
			Percent thermal noise	0.02%	0.02%
			Percent flicker noise	0.00%	0.00%
			Measured transmittance (Tmeas)	0.9977
			Relative transmittance noise (Tnoise)	3.61E-006
			Measured absorbance (Ameas)	0.0010
			Relative standard deviation of absorbance	15.669%
			Absorbance signal-to-noise ratio	6.3820