## Sheet1
Line-source atomic absorption with continuum background correction
Shift, pm | abs. width, pm | Max. Conc. | Background A | Other element | Peak abs. |  | Corrected A
1 | 3 | 1 | 0.1 | 0.2 | 0.33333 |  | 0.18409
| lamp width, pm | stray light | non-abs. line | Hyperfine |  | Cont. Abs | Line Abs
| 1 | 0.1% | 0.1 | 0.1 |  | 0.11135 | 0.29544
| measured I |  | measured I-zero |  | delta I | SNR | Non-lin. error
| 1.0784 |  | 2.1292 |  | 1.0508 | 191.1669 | 0.65%
|  | 3 | 1 | 2
|  | 1 | 1 | 1
Table of Named variables
[15 empty] | Value | Name | Meaning
[15 empty] | 0.33333 | PA | Peak absorbance
[15 empty] | 1 | SH | Shift
[15 empty] | 3 | AW | Absorption width
[15 empty] | 0.2 | IN | Interfering line (Other element)
[15 empty] | 0.1 | HY | hyperfine splitting
[9 empty] | calibration curve | [5 empty] | 0.001 | SL | Stray Light
[9 empty] | slope | 0.24896 |  |  |  |  | 0.1 | BA | Background absorbance
[9 empty] | Intercept | -0.014949 |  |  |  |  | 0.1 | NA | Non-absorbing line from lamp
[9 empty] | Rsquared (R2) | 0.99962 |  |  |  |  | 1 | LW | Line width from lamp
[15 empty] | 1.0784 | MI | Measured transmitted intensity (I)
[15 empty] | 2.1292 | MIZ | Measured I zero
[15 empty] | 0.11135 | CA | Continuum absorbance
[15 empty] | 0.29544 | LA | Line absorbance
© T. C. O'Haver, August 17, 2014 | [8 empty] | 1
|  |  | source | transmitted | Concentrations =>
wavelength | absorbance | transmission | source int | trans int | 0 | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.7 | 0.8 | 0.9 | 1
-100 | 9.4327e-05 | 0.79416 | 1e-05 | 7.9416e-06 | 7.9448e-06 | 8.0666e-06 | 8.0967e-06 | 8.0656e-06 | 8.065e-06 | 8.0645e-06 | 8.0639e-06 | 8.0634e-06 | 8.0628e-06 | 8.0623e-06 | 8.0617e-06
-99 | 9.6095e-05 | 0.79415 | 1e-05 | 7.9415e-06 | 7.9448e-06 | 8.0666e-06 | 8.0967e-06 | 8.0655e-06 | 8.065e-06 | 8.0644e-06 | 8.0639e-06 | 8.0633e-06 | 8.0627e-06 | 8.0622e-06 | 8.0616e-06
-98 | 9.7914e-05 | 0.79415 | 1e-05 | 7.9415e-06 | 7.9448e-06 | 8.0666e-06 | 8.0967e-06 | 8.0655e-06 | 8.0649e-06 | 8.0644e-06 | 8.0638e-06 | 8.0632e-06 | 8.0627e-06 | 8.0621e-06 | 8.0615e-06
-97 | 9.9787e-05 | 0.79415 | 1e-05 | 7.9415e-06 | 7.9448e-06 | 8.0666e-06 | 8.0966e-06 | 8.0655e-06 | 8.0649e-06 | 8.0643e-06 | 8.0637e-06 | 8.0632e-06 | 8.0626e-06 | 8.062e-06 | 8.0614e-06
-96 | 0.00010172 | 0.79414 | 1e-05 | 7.9414e-06 | 7.9448e-06 | 8.0666e-06 | 8.0966e-06 | 8.0655e-06 | 8.0649e-06 | 8.0643e-06 | 8.0637e-06 | 8.0631e-06 | 8.0625e-06 | 8.0619e-06 | 8.0613e-06
-95 | 0.0001037 | 0.79414 | 1e-05 | 7.9414e-06 | 7.9448e-06 | 8.0666e-06 | 8.0966e-06 | 8.0654e-06 | 8.0648e-06 | 8.0642e-06 | 8.0636e-06 | 8.063e-06 | 8.0624e-06 | 8.0618e-06 | 8.0612e-06
-94 | 0.00010575 | 0.79413 | 1e-05 | 7.9413e-06 | 7.9449e-06 | 8.0666e-06 | 8.0966e-06 | 8.0654e-06 | 8.0648e-06 | 8.0642e-06 | 8.0636e-06 | 8.0629e-06 | 8.0623e-06 | 8.0617e-06 | 8.0611e-06
-93 | 0.00010786 | 0.79413 | 1e-05 | 7.9413e-06 | 7.9449e-06 | 8.0666e-06 | 8.0966e-06 | 8.0654e-06 | 8.0648e-06 | 8.0641e-06 | 8.0635e-06 | 8.0629e-06 | 8.0622e-06 | 8.0616e-06 | 8.061e-06
-92 | 0.00011004 | 0.79413 | 1e-05 | 7.9413e-06 | 7.9449e-06 | 8.0666e-06 | 8.0966e-06 | 8.0654e-06 | 8.0647e-06 | 8.0641e-06 | 8.0634e-06 | 8.0628e-06 | 8.0621e-06 | 8.0615e-06 | 8.0609e-06
-91 | 0.00011228 | 0.79412 | 1e-05 | 7.9412e-06 | 7.9449e-06 | 8.0666e-06 | 8.0966e-06 | 8.0653e-06 | 8.0647e-06 | 8.064e-06 | 8.0634e-06 | 8.0627e-06 | 8.062e-06 | 8.0614e-06 | 8.0607e-06
-90 | 0.0001146 | 0.79412 | 1e-05 | 7.9412e-06 | 7.9449e-06 | 8.0666e-06 | 8.0965e-06 | 8.0653e-06 | 8.0646e-06 | 8.064e-06 | 8.0633e-06 | 8.0626e-06 | 8.0619e-06 | 8.0613e-06 | 8.0606e-06
-89 | 0.00011699 | 0.79411 | 1e-05 | 7.9411e-06 | 7.9449e-06 | 8.0666e-06 | 8.0965e-06 | 8.0653e-06 | 8.0646e-06 | 8.0639e-06 | 8.0632e-06 | 8.0625e-06 | 8.0618e-06 | 8.0611e-06 | 8.0605e-06
-88 | 0.00011946 | 0.79411 | 1e-05 | 7.9411e-06 | 7.9449e-06 | 8.0666e-06 | 8.0965e-06 | 8.0652e-06 | 8.0645e-06 | 8.0638e-06 | 8.0631e-06 | 8.0624e-06 | 8.0617e-06 | 8.061e-06 | 8.0603e-06
-87 | 0.00012201 | 0.79411 | 1e-05 | 7.9411e-06 | 7.945e-06 | 8.0666e-06 | 8.0965e-06 | 8.0652e-06 | 8.0645e-06 | 8.0638e-06 | 8.063e-06 | 8.0623e-06 | 8.0616e-06 | 8.0609e-06 | 8.0602e-06
-86 | 0.00012464 | 0.7941 | 1e-05 | 7.941e-06 | 7.945e-06 | 8.0666e-06 | 8.0965e-06 | 8.0652e-06 | 8.0644e-06 | 8.0637e-06 | 8.063e-06 | 8.0622e-06 | 8.0615e-06 | 8.0608e-06 | 8.06e-06
-85 | 0.00012737 | 0.7941 | 1e-05 | 7.941e-06 | 7.945e-06 | 8.0666e-06 | 8.0965e-06 | 8.0651e-06 | 8.0644e-06 | 8.0636e-06 | 8.0629e-06 | 8.0621e-06 | 8.0614e-06 | 8.0606e-06 | 8.0599e-06
-84 | 0.00013019 | 0.79409 | 1e-05 | 7.9409e-06 | 7.945e-06 | 8.0666e-06 | 8.0964e-06 | 8.0651e-06 | 8.0643e-06 | 8.0635e-06 | 8.0628e-06 | 8.062e-06 | 8.0612e-06 | 8.0605e-06 | 8.0597e-06
-83 | 0.0001331 | 0.79408 | 1e-05 | 7.9408e-06 | 7.945e-06 | 8.0666e-06 | 8.0964e-06 | 8.065e-06 | 8.0643e-06 | 8.0635e-06 | 8.0627e-06 | 8.0619e-06 | 8.0611e-06 | 8.0603e-06 | 8.0595e-06
-82 | 0.00013612 | 0.79408 | 1e-05 | 7.9408e-06 | 7.945e-06 | 8.0666e-06 | 8.0964e-06 | 8.065e-06 | 8.0642e-06 | 8.0634e-06 | 8.0626e-06 | 8.0618e-06 | 8.061e-06 | 8.0602e-06 | 8.0594e-06
-81 | 0.00013924 | 0.79407 | 1e-05 | 7.9407e-06 | 7.945e-06 | 8.0666e-06 | 8.0964e-06 | 8.065e-06 | 8.0641e-06 | 8.0633e-06 | 8.0625e-06 | 8.0617e-06 | 8.0608e-06 | 8.06e-06 | 8.0592e-06
-80 | 0.00014248 | 0.79407 | 1e-05 | 7.9407e-06 | 7.9451e-06 | 8.0666e-06 | 8.0963e-06 | 8.0649e-06 | 8.0641e-06 | 8.0632e-06 | 8.0624e-06 | 8.0615e-06 | 8.0607e-06 | 8.0598e-06 | 8.059e-06
-79 | 0.00014583 | 0.79406 | 1e-05 | 7.9406e-06 | 7.9451e-06 | 8.0666e-06 | 8.0963e-06 | 8.0649e-06 | 8.064e-06 | 8.0631e-06 | 8.0623e-06 | 8.0614e-06 | 8.0605e-06 | 8.0597e-06 | 8.0588e-06
-78 | 0.0001493 | 0.79406 | 1e-05 | 7.9406e-06 | 7.9451e-06 | 8.0666e-06 | 8.0963e-06 | 8.0648e-06 | 8.0639e-06 | 8.063e-06 | 8.0621e-06 | 8.0613e-06 | 8.0604e-06 | 8.0595e-06 | 8.0586e-06
-77 | 0.00015291 | 0.79405 | 1e-05 | 7.9405e-06 | 7.9451e-06 | 8.0666e-06 | 8.0963e-06 | 8.0648e-06 | 8.0639e-06 | 8.0629e-06 | 8.062e-06 | 8.0611e-06 | 8.0602e-06 | 8.0593e-06 | 8.0584e-06
-76 | 0.00015665 | 0.79404 | 1e-05 | 7.9404e-06 | 7.9451e-06 | 8.0666e-06 | 8.0962e-06 | 8.0647e-06 | 8.0638e-06 | 8.0628e-06 | 8.0619e-06 | 8.061e-06 | 8.06e-06 | 8.0591e-06 | 8.0582e-06
-75 | 0.00016054 | 0.79403 | 1e-05 | 7.9403e-06 | 7.9451e-06 | 8.0666e-06 | 8.0962e-06 | 8.0647e-06 | 8.0637e-06 | 8.0627e-06 | 8.0618e-06 | 8.0608e-06 | 8.0598e-06 | 8.0589e-06 | 8.0579e-06
-74 | 0.00016458 | 0.79403 | 1e-05 | 7.9403e-06 | 7.9452e-06 | 8.0666e-06 | 8.0962e-06 | 8.0646e-06 | 8.0636e-06 | 8.0626e-06 | 8.0616e-06 | 8.0606e-06 | 8.0597e-06 | 8.0587e-06 | 8.0577e-06
-73 | 0.00016877 | 0.79402 | 1e-05 | 7.9402e-06 | 7.9452e-06 | 8.0666e-06 | 8.0961e-06 | 8.0645e-06 | 8.0635e-06 | 8.0625e-06 | 8.0615e-06 | 8.0605e-06 | 8.0595e-06 | 8.0584e-06 | 8.0574e-06
-72 | 0.00017313 | 0.79401 | 1e-05 | 7.9401e-06 | 7.9452e-06 | 8.0666e-06 | 8.0961e-06 | 8.0645e-06 | 8.0634e-06 | 8.0624e-06 | 8.0613e-06 | 8.0603e-06 | 8.0593e-06 | 8.0582e-06 | 8.0572e-06
-71 | 0.00017768 | 0.794 | 1e-05 | 7.94e-06 | 7.9452e-06 | 8.0666e-06 | 8.0961e-06 | 8.0644e-06 | 8.0633e-06 | 8.0623e-06 | 8.0612e-06 | 8.0601e-06 | 8.059e-06 | 8.058e-06 | 8.0569e-06
-70 | 0.0001824 | 0.79399 | 1e-05 | 7.9399e-06 | 7.9452e-06 | 8.0665e-06 | 8.096e-06 | 8.0643e-06 | 8.0632e-06 | 8.0621e-06 | 8.061e-06 | 8.0599e-06 | 8.0588e-06 | 8.0577e-06 | 8.0566e-06
-69 | 0.00018733 | 0.79399 | 1e-05 | 7.9399e-06 | 7.9453e-06 | 8.0665e-06 | 8.096e-06 | 8.0643e-06 | 8.0631e-06 | 8.062e-06 | 8.0609e-06 | 8.0597e-06 | 8.0586e-06 | 8.0575e-06 | 8.0563e-06
-68 | 0.00019246 | 0.79398 | 1e-05 | 7.9398e-06 | 7.9453e-06 | 8.0665e-06 | 8.0959e-06 | 8.0642e-06 | 8.063e-06 | 8.0618e-06 | 8.0607e-06 | 8.0595e-06 | 8.0583e-06 | 8.0572e-06 | 8.056e-06
-67 | 0.00019781 | 0.79397 | 1e-05 | 7.9397e-06 | 7.9453e-06 | 8.0665e-06 | 8.0959e-06 | 8.0641e-06 | 8.0629e-06 | 8.0617e-06 | 8.0605e-06 | 8.0593e-06 | 8.0581e-06 | 8.0569e-06 | 8.0557e-06
-66 | 0.0002034 | 0.79396 | 1e-05 | 7.9396e-06 | 7.9453e-06 | 8.0665e-06 | 8.0958e-06 | 8.064e-06 | 8.0628e-06 | 8.0615e-06 | 8.0603e-06 | 8.0591e-06 | 8.0578e-06 | 8.0566e-06 | 8.0553e-06
-65 | 0.00020923 | 0.79395 | 1e-05 | 7.9395e-06 | 7.9453e-06 | 8.0665e-06 | 8.0958e-06 | 8.0639e-06 | 8.0626e-06 | 8.0614e-06 | 8.0601e-06 | 8.0588e-06 | 8.0575e-06 | 8.0563e-06 | 8.055e-06
-64 | 0.00021533 | 0.79393 | 1.0011e-05 | 7.9483e-06 | 7.9543e-06 | 8.0755e-06 | 8.1048e-06 | 8.0729e-06 | 8.0716e-06 | 8.0703e-06 | 8.0689e-06 | 8.0676e-06 | 8.0663e-06 | 8.065e-06 | 8.0637e-06
-63 | 0.00022171 | 0.79392 | 2.2341e-05 | 1.7737e-05 | 1.7751e-05 | 1.8021e-05 | 1.8086e-05 | 1.8015e-05 | 1.8012e-05 | 1.8009e-05 | 1.8006e-05 | 1.8003e-05 | 1.8e-05 | 1.7997e-05 | 1.7994e-05
-62 | 0.00022838 | 0.79391 | 0.0018416 | 0.001462 | 0.0014632 | 0.0014855 | 0.0014909 | 0.001485 | 0.0014847 | 0.0014845 | 0.0014842 | 0.0014839 | 0.0014837 | 0.0014834 | 0.0014832
-61 | 0.00023537 | 0.7939 | 0.036798 | 0.029214 | 0.029238 | 0.029683 | 0.02979 | 0.029672 | 0.029667 | 0.029661 | 0.029656 | 0.029651 | 0.029645 | 0.02964 | 0.029635
-60 | 0.00024269 | 0.79388 | 0.10001 | 0.079396 | 0.079462 | 0.080672 | 0.080963 | 0.080642 | 0.080627 | 0.080612 | 0.080597 | 0.080582 | 0.080567 | 0.080552 | 0.080537
-59 | 0.00025038 | 0.79387 | 0.036798 | 0.029213 | 0.029238 | 0.029682 | 0.029789 | 0.029671 | 0.029665 | 0.02966 | 0.029654 | 0.029648 | 0.029643 | 0.029637 | 0.029631
-58 | 0.00025844 | 0.79386 | 0.0018416 | 0.0014619 | 0.0014632 | 0.0014855 | 0.0014908 | 0.0014849 | 0.0014846 | 0.0014843 | 0.001484 | 0.0014837 | 0.0014834 | 0.0014831 | 0.0014828
-57 | 0.00026692 | 0.79384 | 2.2341e-05 | 1.7735e-05 | 1.7751e-05 | 1.8021e-05 | 1.8085e-05 | 1.8013e-05 | 1.801e-05 | 1.8006e-05 | 1.8002e-05 | 1.7999e-05 | 1.7995e-05 | 1.7991e-05 | 1.7988e-05
-56 | 0.00027583 | 0.79382 | 1.0011e-05 | 7.9472e-06 | 7.9545e-06 | 8.0753e-06 | 8.1042e-06 | 8.0719e-06 | 8.0702e-06 | 8.0685e-06 | 8.0668e-06 | 8.0651e-06 | 8.0634e-06 | 8.0616e-06 | 8.0599e-06
-55 | 0.00028521 | 0.79381 | 1e-05 | 7.9381e-06 | 7.9456e-06 | 8.0662e-06 | 8.095e-06 | 8.0627e-06 | 8.0609e-06 | 8.0591e-06 | 8.0574e-06 | 8.0556e-06 | 8.0538e-06 | 8.0521e-06 | 8.0503e-06
-54 | 0.00029509 | 0.79379 | 1e-05 | 7.9379e-06 | 7.9456e-06 | 8.0662e-06 | 8.0949e-06 | 8.0625e-06 | 8.0607e-06 | 8.0588e-06 | 8.057e-06 | 8.0552e-06 | 8.0533e-06 | 8.0515e-06 | 8.0497e-06
-53 | 0.00030551 | 0.79377 | 1e-05 | 7.9377e-06 | 7.9456e-06 | 8.0661e-06 | 8.0948e-06 | 8.0623e-06 | 8.0604e-06 | 8.0585e-06 | 8.0566e-06 | 8.0547e-06 | 8.0528e-06 | 8.0509e-06 | 8.049e-06
-52 | 0.00031651 | 0.79375 | 1e-05 | 7.9375e-06 | 7.9457e-06 | 8.0661e-06 | 8.0947e-06 | 8.0621e-06 | 8.0602e-06 | 8.0582e-06 | 8.0562e-06 | 8.0542e-06 | 8.0523e-06 | 8.0503e-06 | 8.0483e-06
-51 | 0.00032813 | 0.79373 | 1e-05 | 7.9373e-06 | 7.9457e-06 | 8.0661e-06 | 8.0946e-06 | 8.0619e-06 | 8.0599e-06 | 8.0578e-06 | 8.0558e-06 | 8.0537e-06 | 8.0517e-06 | 8.0496e-06 | 8.0476e-06
-50 | 0.00034043 | 0.79371 | 1e-05 | 7.9371e-06 | 7.9457e-06 | 8.066e-06 | 8.0944e-06 | 8.0617e-06 | 8.0596e-06 | 8.0575e-06 | 8.0553e-06 | 8.0532e-06 | 8.0511e-06 | 8.0489e-06 | 8.0468e-06
-49 | 0.00035345 | 0.79368 | 1e-05 | 7.9368e-06 | 7.9457e-06 | 8.0659e-06 | 8.0943e-06 | 8.0615e-06 | 8.0593e-06 | 8.0571e-06 | 8.0548e-06 | 8.0526e-06 | 8.0504e-06 | 8.0482e-06 | 8.046e-06
-48 | 0.00036725 | 0.79366 | 1e-05 | 7.9366e-06 | 7.9458e-06 | 8.0659e-06 | 8.0941e-06 | 8.0613e-06 | 8.059e-06 | 8.0566e-06 | 8.0543e-06 | 8.052e-06 | 8.0497e-06 | 8.0474e-06 | 8.0451e-06
-47 | 0.00038191 | 0.79363 | 1e-05 | 7.9363e-06 | 7.9458e-06 | 8.0658e-06 | 8.094e-06 | 8.061e-06 | 8.0586e-06 | 8.0562e-06 | 8.0538e-06 | 8.0514e-06 | 8.049e-06 | 8.0466e-06 | 8.0442e-06
-46 | 0.00039749 | 0.7936 | 1e-05 | 7.936e-06 | 7.9458e-06 | 8.0658e-06 | 8.0938e-06 | 8.0607e-06 | 8.0582e-06 | 8.0557e-06 | 8.0532e-06 | 8.0507e-06 | 8.0482e-06 | 8.0457e-06 | 8.0432e-06
-45 | 0.00041407 | 0.79357 | 1e-05 | 7.9357e-06 | 7.9459e-06 | 8.0657e-06 | 8.0936e-06 | 8.0604e-06 | 8.0578e-06 | 8.0552e-06 | 8.0526e-06 | 8.05e-06 | 8.0473e-06 | 8.0447e-06 | 8.0421e-06
-44 | 0.00043174 | 0.79354 | 1e-05 | 7.9354e-06 | 7.9459e-06 | 8.0656e-06 | 8.0934e-06 | 8.0601e-06 | 8.0574e-06 | 8.0546e-06 | 8.0519e-06 | 8.0492e-06 | 8.0464e-06 | 8.0437e-06 | 8.041e-06
-43 | 0.0004506 | 0.7935 | 1e-05 | 7.935e-06 | 7.9459e-06 | 8.0655e-06 | 8.0932e-06 | 8.0598e-06 | 8.0569e-06 | 8.0541e-06 | 8.0512e-06 | 8.0483e-06 | 8.0455e-06 | 8.0426e-06 | 8.0398e-06
-42 | 0.00047077 | 0.79347 | 1e-05 | 7.9347e-06 | 7.946e-06 | 8.0654e-06 | 8.093e-06 | 8.0594e-06 | 8.0564e-06 | 8.0534e-06 | 8.0504e-06 | 8.0474e-06 | 8.0444e-06 | 8.0415e-06 | 8.0385e-06
-41 | 0.00049236 | 0.79343 | 1e-05 | 7.9343e-06 | 7.946e-06 | 8.0653e-06 | 8.0927e-06 | 8.059e-06 | 8.0559e-06 | 8.0527e-06 | 8.0496e-06 | 8.0465e-06 | 8.0433e-06 | 8.0402e-06 | 8.0371e-06
-40 | 0.00051551 | 0.79339 | 1e-05 | 7.9339e-06 | 7.9461e-06 | 8.0652e-06 | 8.0924e-06 | 8.0586e-06 | 8.0553e-06 | 8.052e-06 | 8.0487e-06 | 8.0454e-06 | 8.0421e-06 | 8.0389e-06 | 8.0356e-06
-39 | 0.00054038 | 0.79334 | 1e-05 | 7.9334e-06 | 7.9461e-06 | 8.0651e-06 | 8.0922e-06 | 8.0581e-06 | 8.0547e-06 | 8.0512e-06 | 8.0478e-06 | 8.0443e-06 | 8.0409e-06 | 8.0374e-06 | 8.034e-06
-38 | 0.00056714 | 0.79329 | 1e-05 | 7.9329e-06 | 7.9461e-06 | 8.0649e-06 | 8.0918e-06 | 8.0576e-06 | 8.054e-06 | 8.0504e-06 | 8.0467e-06 | 8.0431e-06 | 8.0395e-06 | 8.0358e-06 | 8.0322e-06
-37 | 0.00059601 | 0.79324 | 1e-05 | 7.9324e-06 | 7.9462e-06 | 8.0648e-06 | 8.0915e-06 | 8.0571e-06 | 8.0533e-06 | 8.0494e-06 | 8.0456e-06 | 8.0418e-06 | 8.038e-06 | 8.0341e-06 | 8.0303e-06
-36 | 0.00062719 | 0.79318 | 1e-05 | 7.9318e-06 | 7.9462e-06 | 8.0646e-06 | 8.0911e-06 | 8.0565e-06 | 8.0525e-06 | 8.0484e-06 | 8.0444e-06 | 8.0404e-06 | 8.0363e-06 | 8.0323e-06 | 8.0283e-06
-35 | 0.00066096 | 0.79312 | 1e-05 | 7.9312e-06 | 7.9463e-06 | 8.0644e-06 | 8.0907e-06 | 8.0559e-06 | 8.0516e-06 | 8.0474e-06 | 8.0431e-06 | 8.0388e-06 | 8.0346e-06 | 8.0303e-06 | 8.0261e-06
-34 | 0.0006976 | 0.79305 | 1e-05 | 7.9305e-06 | 7.9463e-06 | 8.0642e-06 | 8.0903e-06 | 8.0552e-06 | 8.0507e-06 | 8.0462e-06 | 8.0417e-06 | 8.0372e-06 | 8.0327e-06 | 8.0282e-06 | 8.0237e-06
-33 | 0.00073746 | 0.79298 | 1e-05 | 7.9298e-06 | 7.9463e-06 | 8.064e-06 | 8.0898e-06 | 8.0544e-06 | 8.0497e-06 | 8.0449e-06 | 8.0401e-06 | 8.0354e-06 | 8.0306e-06 | 8.0258e-06 | 8.0211e-06
-32 | 0.00078092 | 0.7929 | 1e-05 | 7.929e-06 | 7.9464e-06 | 8.0637e-06 | 8.0892e-06 | 8.0536e-06 | 8.0485e-06 | 8.0435e-06 | 8.0384e-06 | 8.0334e-06 | 8.0283e-06 | 8.0233e-06 | 8.0182e-06
-31 | 0.00082844 | 0.79281 | 1e-05 | 7.9281e-06 | 7.9464e-06 | 8.0635e-06 | 8.0886e-06 | 8.0527e-06 | 8.0473e-06 | 8.0419e-06 | 8.0366e-06 | 8.0312e-06 | 8.0258e-06 | 8.0205e-06 | 8.0151e-06
-30 | 0.00088055 | 0.79272 | 1e-05 | 7.9272e-06 | 7.9465e-06 | 8.0632e-06 | 8.088e-06 | 8.0517e-06 | 8.046e-06 | 8.0402e-06 | 8.0345e-06 | 8.0288e-06 | 8.0231e-06 | 8.0174e-06 | 8.0117e-06
-29 | 0.00093785 | 0.79261 | 1e-05 | 7.9261e-06 | 7.9465e-06 | 8.0628e-06 | 8.0873e-06 | 8.0506e-06 | 8.0445e-06 | 8.0384e-06 | 8.0323e-06 | 8.0262e-06 | 8.0201e-06 | 8.014e-06 | 8.0079e-06
-28 | 0.0010011 | 0.7925 | 1e-05 | 7.925e-06 | 7.9466e-06 | 8.0625e-06 | 8.0865e-06 | 8.0494e-06 | 8.0428e-06 | 8.0363e-06 | 8.0298e-06 | 8.0232e-06 | 8.0167e-06 | 8.0102e-06 | 8.0037e-06
-27 | 0.0010711 | 0.79237 | 1e-05 | 7.9237e-06 | 7.9466e-06 | 8.0621e-06 | 8.0856e-06 | 8.048e-06 | 8.041e-06 | 8.034e-06 | 8.027e-06 | 8.02e-06 | 8.013e-06 | 8.0061e-06 | 7.9991e-06
-26 | 0.0011488 | 0.79223 | 1e-05 | 7.9223e-06 | 7.9467e-06 | 8.0616e-06 | 8.0846e-06 | 8.0465e-06 | 8.039e-06 | 8.0315e-06 | 8.0239e-06 | 8.0164e-06 | 8.0089e-06 | 8.0014e-06 | 7.994e-06
-25 | 0.0012355 | 0.79207 | 1e-05 | 7.9207e-06 | 7.9468e-06 | 8.0611e-06 | 8.0835e-06 | 8.0448e-06 | 8.0367e-06 | 8.0286e-06 | 8.0205e-06 | 8.0124e-06 | 8.0044e-06 | 7.9963e-06 | 7.9882e-06
-24 | 0.0013327 | 0.79189 | 1e-05 | 7.9189e-06 | 7.9468e-06 | 8.0605e-06 | 8.0822e-06 | 8.0429e-06 | 8.0342e-06 | 8.0254e-06 | 8.0167e-06 | 8.0079e-06 | 7.9992e-06 | 7.9905e-06 | 7.9818e-06
-23 | 0.001442 | 0.7917 | 1e-05 | 7.917e-06 | 7.9469e-06 | 8.0598e-06 | 8.0808e-06 | 8.0408e-06 | 8.0313e-06 | 8.0218e-06 | 8.0123e-06 | 8.0029e-06 | 7.9934e-06 | 7.984e-06 | 7.9746e-06
-22 | 0.0015656 | 0.79147 | 1e-05 | 7.9147e-06 | 7.9469e-06 | 8.059e-06 | 8.0792e-06 | 8.0383e-06 | 8.028e-06 | 8.0177e-06 | 8.0074e-06 | 7.9972e-06 | 7.9869e-06 | 7.9767e-06 | 7.9664e-06
-21 | 0.0017061 | 0.79121 | 1e-05 | 7.9121e-06 | 7.947e-06 | 8.0581e-06 | 8.0774e-06 | 8.0356e-06 | 8.0243e-06 | 8.0131e-06 | 8.0019e-06 | 7.9907e-06 | 7.9795e-06 | 7.9683e-06 | 7.9572e-06
-20 | 0.0018667 | 0.79092 | 1e-05 | 7.9092e-06 | 7.9471e-06 | 8.0571e-06 | 8.0753e-06 | 8.0324e-06 | 8.0201e-06 | 8.0078e-06 | 7.9955e-06 | 7.9833e-06 | 7.971e-06 | 7.9588e-06 | 7.9466e-06
-19 | 0.0020516 | 0.79058 | 1e-05 | 7.9058e-06 | 7.9471e-06 | 8.056e-06 | 8.0728e-06 | 8.0288e-06 | 8.0152e-06 | 8.0017e-06 | 7.9882e-06 | 7.9747e-06 | 7.9613e-06 | 7.9478e-06 | 7.9344e-06
-18 | 0.0022658 | 0.79019 | 1e-05 | 7.9019e-06 | 7.9472e-06 | 8.0546e-06 | 8.07e-06 | 8.0246e-06 | 8.0096e-06 | 7.9947e-06 | 7.9797e-06 | 7.9649e-06 | 7.95e-06 | 7.9352e-06 | 7.9204e-06
-17 | 0.0025159 | 0.78974 | 1e-05 | 7.8974e-06 | 7.9473e-06 | 8.053e-06 | 8.0668e-06 | 8.0196e-06 | 8.003e-06 | 7.9864e-06 | 7.9699e-06 | 7.9534e-06 | 7.9369e-06 | 7.9204e-06 | 7.904e-06
-16 | 0.0028104 | 0.7892 | 1e-05 | 7.892e-06 | 7.9473e-06 | 8.0511e-06 | 8.0629e-06 | 8.0139e-06 | 7.9953e-06 | 7.9768e-06 | 7.9583e-06 | 7.9399e-06 | 7.9215e-06 | 7.9032e-06 | 7.8849e-06
-15 | 0.0031605 | 0.78857 | 1e-05 | 7.8857e-06 | 7.9474e-06 | 8.0489e-06 | 8.0583e-06 | 8.007e-06 | 7.9862e-06 | 7.9654e-06 | 7.9446e-06 | 7.9239e-06 | 7.9033e-06 | 7.8827e-06 | 7.8622e-06
-14 | 0.0035811 | 0.78781 | 1e-05 | 7.8781e-06 | 7.9475e-06 | 8.0462e-06 | 8.0528e-06 | 7.9988e-06 | 7.9752e-06 | 7.9517e-06 | 7.9283e-06 | 7.9049e-06 | 7.8816e-06 | 7.8584e-06 | 7.8352e-06
-13 | 0.0040926 | 0.78688 | 1e-05 | 7.8688e-06 | 7.9476e-06 | 8.0429e-06 | 8.0462e-06 | 7.9889e-06 | 7.962e-06 | 7.9352e-06 | 7.9085e-06 | 7.8819e-06 | 7.8554e-06 | 7.829e-06 | 7.8026e-06
-12 | 0.0047231 | 0.78574 | 1e-05 | 7.8574e-06 | 7.9477e-06 | 8.0388e-06 | 8.0381e-06 | 7.9767e-06 | 7.9458e-06 | 7.915e-06 | 7.8844e-06 | 7.8538e-06 | 7.8234e-06 | 7.7931e-06 | 7.7629e-06
-11 | 0.0055126 | 0.78431 | 1e-05 | 7.8431e-06 | 7.9477e-06 | 8.0338e-06 | 8.028e-06 | 7.9616e-06 | 7.9258e-06 | 7.8901e-06 | 7.8545e-06 | 7.8191e-06 | 7.7839e-06 | 7.7488e-06 | 7.7139e-06
-10 | 0.006519 | 0.78249 | 1e-05 | 7.8249e-06 | 7.9478e-06 | 8.0275e-06 | 8.0152e-06 | 7.9427e-06 | 7.9006e-06 | 7.8587e-06 | 7.817e-06 | 7.7756e-06 | 7.7344e-06 | 7.6934e-06 | 7.6526e-06
-9 | 0.0078297 | 0.78014 | 1e-05 | 7.8014e-06 | 7.9479e-06 | 8.0194e-06 | 7.9989e-06 | 7.9184e-06 | 7.8684e-06 | 7.8187e-06 | 7.7693e-06 | 7.7202e-06 | 7.6714e-06 | 7.6229e-06 | 7.5748e-06
-8 | 0.0095799 | 0.777 | 1e-05 | 7.77e-06 | 7.948e-06 | 8.0088e-06 | 7.9777e-06 | 7.8868e-06 | 7.8265e-06 | 7.7666e-06 | 7.7072e-06 | 7.6483e-06 | 7.5898e-06 | 7.5318e-06 | 7.4742e-06
-7 | 0.011989 | 0.7727 | 1e-05 | 7.727e-06 | 7.9481e-06 | 7.9946e-06 | 7.9493e-06 | 7.8447e-06 | 7.7708e-06 | 7.6976e-06 | 7.6251e-06 | 7.5533e-06 | 7.4822e-06 | 7.4117e-06 | 7.3419e-06
-6 | 0.015428 | 0.7666 | 1e-05 | 7.666e-06 | 7.9482e-06 | 7.9751e-06 | 7.9106e-06 | 7.7874e-06 | 7.6952e-06 | 7.6041e-06 | 7.5141e-06 | 7.4251e-06 | 7.3372e-06 | 7.2503e-06 | 7.1645e-06
-5 | 0.020571 | 0.75758 | 1e-05 | 7.5758e-06 | 7.9483e-06 | 7.9478e-06 | 7.8564e-06 | 7.7075e-06 | 7.5901e-06 | 7.4745e-06 | 7.3606e-06 | 7.2485e-06 | 7.1381e-06 | 7.0294e-06 | 6.9223e-06
-4 | 0.028719 | 0.7435 | 1.0113e-05 | 7.5187e-06 | 8.0379e-06 | 7.9976e-06 | 7.8664e-06 | 7.6791e-06 | 7.5247e-06 | 7.3733e-06 | 7.225e-06 | 7.0797e-06 | 6.9373e-06 | 6.7978e-06 | 6.6611e-06
-3 | 0.042625 | 0.72007 | 0.00013341 | 9.6065e-05 | 0.00010604 | 0.00010474 | 0.00010227 | 9.911e-05 | 9.6409e-05 | 9.3781e-05 | 9.1225e-05 | 8.8739e-05 | 8.6321e-05 | 8.3968e-05 | 8.168e-05
-2 | 0.068694 | 0.67812 | 0.018326 | 0.012427 | 0.014566 | 0.014234 | 0.01375 | 0.013182 | 0.012686 | 0.012208 | 0.011749 | 0.011306 | 0.010881 | 0.010471 | 0.010077
-1 | 0.12282 | 0.59866 | 0.36789 | 0.22024 | 0.29243 | 0.28157 | 0.26801 | 0.25318 | 0.24008 | 0.22765 | 0.21587 | 0.2047 | 0.19411 | 0.18406 | 0.17454
0 | 0.23495 | 0.46244 | 1 | 0.46245 | 0.79491 | 0.75328 | 0.70567 | 0.65608 | 0.61229 | 0.57142 | 0.53328 | 0.49769 | 0.46447 | 0.43347 | 0.40454
1 | 0.34007 | 0.36302 | 0.36972 | 0.13422 | 0.29389 | 0.27637 | 0.25692 | 0.23704 | 0.21953 | 0.20331 | 0.18829 | 0.17438 | 0.1615 | 0.14956 | 0.13851
2 | 0.24284 | 0.45411 | 0.055114 | 0.025027 | 0.043811 | 0.041517 | 0.038892 | 0.036159 | 0.033746 | 0.031493 | 0.029391 | 0.02743 | 0.025599 | 0.02389 | 0.022296
3 | 0.14315 | 0.57128 | 0.10013 | 0.057204 | 0.079599 | 0.076643 | 0.072952 | 0.068916 | 0.065349 | 0.061968 | 0.058761 | 0.05572 | 0.052836 | 0.050102 | 0.047509
4 | 0.10008 | 0.63084 | 0.036798 | 0.023214 | 0.029252 | 0.028585 | 0.027613 | 0.026473 | 0.025476 | 0.024517 | 0.023594 | 0.022706 | 0.021851 | 0.021028 | 0.020237
5 | 0.064253 | 0.68509 | 0.0018416 | 0.0012616 | 0.001464 | 0.001446 | 0.0014119 | 0.0013683 | 0.001331 | 0.0012947 | 0.0012594 | 0.0012251 | 0.0011917 | 0.0011592 | 0.0011276
6 | 0.039605 | 0.7251 | 2.2341e-05 | 1.6199e-05 | 1.7761e-05 | 1.7672e-05 | 1.7382e-05 | 1.6968e-05 | 1.6627e-05 | 1.6292e-05 | 1.5965e-05 | 1.5643e-05 | 1.5329e-05 | 1.5021e-05 | 1.4718e-05
7 | 0.026359 | 0.74755 | 1.0011e-05 | 7.4839e-06 | 7.9589e-06 | 7.9584e-06 | 7.8669e-06 | 7.7178e-06 | 7.6002e-06 | 7.4844e-06 | 7.3704e-06 | 7.2581e-06 | 7.1476e-06 | 7.0387e-06 | 6.9315e-06
8 | 0.018831 | 0.76062 | 1e-05 | 7.6062e-06 | 7.9501e-06 | 7.977e-06 | 7.9125e-06 | 7.7893e-06 | 7.697e-06 | 7.6059e-06 | 7.5159e-06 | 7.4269e-06 | 7.339e-06 | 7.2521e-06 | 7.1662e-06
9 | 0.014162 | 0.76884 | 1e-05 | 7.6884e-06 | 7.9503e-06 | 7.9968e-06 | 7.9515e-06 | 7.8468e-06 | 7.7729e-06 | 7.6997e-06 | 7.6272e-06 | 7.5554e-06 | 7.4842e-06 | 7.4137e-06 | 7.3439e-06
10 | 0.011062 | 0.77435 | 1e-05 | 7.7435e-06 | 7.9505e-06 | 8.0112e-06 | 7.9801e-06 | 7.8892e-06 | 7.8289e-06 | 7.769e-06 | 7.7096e-06 | 7.6507e-06 | 7.5922e-06 | 7.5341e-06 | 7.4765e-06
11 | 0.0088929 | 0.77823 | 1e-05 | 7.7823e-06 | 7.9507e-06 | 8.0222e-06 | 8.0017e-06 | 7.9211e-06 | 7.8711e-06 | 7.8214e-06 | 7.772e-06 | 7.7229e-06 | 7.6741e-06 | 7.6256e-06 | 7.5774e-06
12 | 0.0073145 | 0.78106 | 1e-05 | 7.8106e-06 | 7.9509e-06 | 8.0306e-06 | 8.0183e-06 | 7.9457e-06 | 7.9036e-06 | 7.8617e-06 | 7.82e-06 | 7.7786e-06 | 7.7374e-06 | 7.6964e-06 | 7.6556e-06
13 | 0.0061293 | 0.7832 | 1e-05 | 7.832e-06 | 7.9511e-06 | 8.0372e-06 | 8.0314e-06 | 7.965e-06 | 7.9291e-06 | 7.8934e-06 | 7.8578e-06 | 7.8224e-06 | 7.7872e-06 | 7.7521e-06 | 7.7172e-06
14 | 0.0052163 | 0.78484 | 1e-05 | 7.8484e-06 | 7.9513e-06 | 8.0426e-06 | 8.0418e-06 | 7.9804e-06 | 7.9495e-06 | 7.9187e-06 | 7.888e-06 | 7.8574e-06 | 7.827e-06 | 7.7967e-06 | 7.7665e-06
15 | 0.0044981 | 0.78614 | 1e-05 | 7.8614e-06 | 7.9516e-06 | 8.0469e-06 | 8.0502e-06 | 7.9928e-06 | 7.966e-06 | 7.9392e-06 | 7.9125e-06 | 7.8858e-06 | 7.8593e-06 | 7.8329e-06 | 7.8065e-06
16 | 0.003923 | 0.78719 | 1e-05 | 7.8719e-06 | 7.9518e-06 | 8.0505e-06 | 8.0572e-06 | 8.0031e-06 | 7.9795e-06 | 7.956e-06 | 7.9326e-06 | 7.9092e-06 | 7.8859e-06 | 7.8626e-06 | 7.8394e-06
17 | 0.0034554 | 0.78803 | 1e-05 | 7.8803e-06 | 7.9521e-06 | 8.0535e-06 | 8.063e-06 | 8.0117e-06 | 7.9908e-06 | 7.97e-06 | 7.9493e-06 | 7.9286e-06 | 7.9079e-06 | 7.8873e-06 | 7.8668e-06
18 | 0.0030704 | 0.78873 | 1e-05 | 7.8873e-06 | 7.9523e-06 | 8.0561e-06 | 8.068e-06 | 8.0189e-06 | 8.0003e-06 | 7.9818e-06 | 7.9633e-06 | 7.9449e-06 | 7.9265e-06 | 7.9081e-06 | 7.8898e-06
19 | 0.0027497 | 0.78931 | 1e-05 | 7.8931e-06 | 7.9526e-06 | 8.0584e-06 | 8.0722e-06 | 8.025e-06 | 8.0084e-06 | 7.9918e-06 | 7.9752e-06 | 7.9587e-06 | 7.9422e-06 | 7.9258e-06 | 7.9093e-06
20 | 0.00248 | 0.78981 | 1e-05 | 7.8981e-06 | 7.9529e-06 | 8.0604e-06 | 8.0758e-06 | 8.0303e-06 | 8.0153e-06 | 8.0004e-06 | 7.9855e-06 | 7.9706e-06 | 7.9557e-06 | 7.9409e-06 | 7.9261e-06
21 | 0.0022512 | 0.79022 | 1e-05 | 7.9022e-06 | 7.9532e-06 | 8.0621e-06 | 8.079e-06 | 8.0349e-06 | 8.0214e-06 | 8.0078e-06 | 7.9943e-06 | 7.9808e-06 | 7.9674e-06 | 7.9539e-06 | 7.9405e-06
22 | 0.0020556 | 0.79058 | 1e-05 | 7.9058e-06 | 7.9535e-06 | 8.0637e-06 | 8.0818e-06 | 8.039e-06 | 8.0266e-06 | 8.0143e-06 | 8.002e-06 | 7.9898e-06 | 7.9775e-06 | 7.9653e-06 | 7.953e-06
23 | 0.0018873 | 0.79088 | 1e-05 | 7.9088e-06 | 7.9539e-06 | 8.0651e-06 | 8.0844e-06 | 8.0425e-06 | 8.0313e-06 | 8.02e-06 | 8.0088e-06 | 7.9976e-06 | 7.9864e-06 | 7.9752e-06 | 7.964e-06
24 | 0.0017416 | 0.79115 | 1e-05 | 7.9115e-06 | 7.9542e-06 | 8.0664e-06 | 8.0866e-06 | 8.0457e-06 | 8.0354e-06 | 8.0251e-06 | 8.0148e-06 | 8.0045e-06 | 7.9943e-06 | 7.984e-06 | 7.9738e-06
25 | 0.0016149 | 0.79138 | 1e-05 | 7.9138e-06 | 7.9546e-06 | 8.0677e-06 | 8.0887e-06 | 8.0486e-06 | 8.0391e-06 | 8.0296e-06 | 8.0202e-06 | 8.0107e-06 | 8.0012e-06 | 7.9918e-06 | 7.9824e-06
26 | 0.0015041 | 0.79158 | 1e-05 | 7.9158e-06 | 7.955e-06 | 8.0688e-06 | 8.0906e-06 | 8.0512e-06 | 8.0425e-06 | 8.0337e-06 | 8.025e-06 | 8.0162e-06 | 8.0075e-06 | 7.9988e-06 | 7.9901e-06
27 | 0.001407 | 0.79176 | 1e-05 | 7.9176e-06 | 7.9555e-06 | 8.0699e-06 | 8.0923e-06 | 8.0536e-06 | 8.0455e-06 | 8.0374e-06 | 8.0293e-06 | 8.0212e-06 | 8.0131e-06 | 8.005e-06 | 7.997e-06
28 | 0.0013216 | 0.79191 | 1e-05 | 7.9191e-06 | 7.9559e-06 | 8.0709e-06 | 8.094e-06 | 8.0558e-06 | 8.0483e-06 | 8.0408e-06 | 8.0332e-06 | 8.0257e-06 | 8.0182e-06 | 8.0107e-06 | 8.0032e-06
29 | 0.0012463 | 0.79205 | 1e-05 | 7.9205e-06 | 7.9564e-06 | 8.0719e-06 | 8.0955e-06 | 8.0579e-06 | 8.0509e-06 | 8.0439e-06 | 8.0369e-06 | 8.0299e-06 | 8.0229e-06 | 8.0159e-06 | 8.0089e-06
30 | 0.0011798 | 0.79217 | 1e-05 | 7.9217e-06 | 7.9569e-06 | 8.0729e-06 | 8.097e-06 | 8.0598e-06 | 8.0533e-06 | 8.0467e-06 | 8.0402e-06 | 8.0337e-06 | 8.0271e-06 | 8.0206e-06 | 8.0141e-06
31 | 0.0011211 | 0.79228 | 1e-05 | 7.9228e-06 | 7.9574e-06 | 8.0739e-06 | 8.0984e-06 | 8.0616e-06 | 8.0555e-06 | 8.0494e-06 | 8.0433e-06 | 8.0372e-06 | 8.0311e-06 | 8.025e-06 | 8.0189e-06
32 | 0.0010691 | 0.79238 | 1e-05 | 7.9238e-06 | 7.958e-06 | 8.0749e-06 | 8.0997e-06 | 8.0634e-06 | 8.0576e-06 | 8.0519e-06 | 8.0462e-06 | 8.0404e-06 | 8.0347e-06 | 8.029e-06 | 8.0233e-06
33 | 0.0010233 | 0.79246 | 1e-05 | 7.9246e-06 | 7.9586e-06 | 8.0758e-06 | 8.101e-06 | 8.065e-06 | 8.0597e-06 | 8.0543e-06 | 8.0489e-06 | 8.0435e-06 | 8.0381e-06 | 8.0328e-06 | 8.0274e-06
34 | 0.00098285 | 0.79253 | 1e-05 | 7.9253e-06 | 7.9593e-06 | 8.0768e-06 | 8.1023e-06 | 8.0667e-06 | 8.0616e-06 | 8.0565e-06 | 8.0515e-06 | 8.0464e-06 | 8.0413e-06 | 8.0363e-06 | 8.0312e-06
35 | 0.00094737 | 0.7926 | 1e-05 | 7.926e-06 | 7.96e-06 | 8.0778e-06 | 8.1036e-06 | 8.0682e-06 | 8.0635e-06 | 8.0587e-06 | 8.0539e-06 | 8.0491e-06 | 8.0444e-06 | 8.0396e-06 | 8.0348e-06
36 | 0.00091637 | 0.79265 | 1e-05 | 7.9265e-06 | 7.9607e-06 | 8.0788e-06 | 8.1049e-06 | 8.0698e-06 | 8.0653e-06 | 8.0608e-06 | 8.0563e-06 | 8.0518e-06 | 8.0472e-06 | 8.0427e-06 | 8.0382e-06
37 | 0.00088951 | 0.7927 | 1e-05 | 7.927e-06 | 7.9615e-06 | 8.0799e-06 | 8.1062e-06 | 8.0713e-06 | 8.0671e-06 | 8.0628e-06 | 8.0585e-06 | 8.0543e-06 | 8.05e-06 | 8.0457e-06 | 8.0415e-06
38 | 0.00086648 | 0.79275 | 1e-05 | 7.9275e-06 | 7.9624e-06 | 8.081e-06 | 8.1076e-06 | 8.0729e-06 | 8.0688e-06 | 8.0648e-06 | 8.0608e-06 | 8.0567e-06 | 8.0527e-06 | 8.0486e-06 | 8.0446e-06
39 | 0.00084703 | 0.79278 | 1e-05 | 7.9278e-06 | 7.9633e-06 | 8.0821e-06 | 8.1089e-06 | 8.0744e-06 | 8.0706e-06 | 8.0668e-06 | 8.0629e-06 | 8.0591e-06 | 8.0553e-06 | 8.0514e-06 | 8.0476e-06
40 | 0.00083099 | 0.79281 | 1e-05 | 7.9281e-06 | 7.9643e-06 | 8.0833e-06 | 8.1103e-06 | 8.076e-06 | 8.0724e-06 | 8.0687e-06 | 8.0651e-06 | 8.0614e-06 | 8.0578e-06 | 8.0542e-06 | 8.0505e-06
41 | 0.00081821 | 0.79283 | 1e-05 | 7.9283e-06 | 7.9653e-06 | 8.0846e-06 | 8.1117e-06 | 8.0776e-06 | 8.0742e-06 | 8.0707e-06 | 8.0672e-06 | 8.0638e-06 | 8.0603e-06 | 8.0568e-06 | 8.0534e-06
42 | 0.00080857 | 0.79285 | 1e-05 | 7.9285e-06 | 7.9665e-06 | 8.0859e-06 | 8.1132e-06 | 8.0793e-06 | 8.076e-06 | 8.0727e-06 | 8.0694e-06 | 8.0661e-06 | 8.0628e-06 | 8.0595e-06 | 8.0562e-06
43 | 0.00080203 | 0.79286 | 1e-05 | 7.9286e-06 | 7.9677e-06 | 8.0873e-06 | 8.1148e-06 | 8.081e-06 | 8.0779e-06 | 8.0747e-06 | 8.0716e-06 | 8.0684e-06 | 8.0653e-06 | 8.0621e-06 | 8.059e-06
44 | 0.00079856 | 0.79287 | 1e-05 | 7.9287e-06 | 7.969e-06 | 8.0888e-06 | 8.1165e-06 | 8.0828e-06 | 8.0798e-06 | 8.0768e-06 | 8.0738e-06 | 8.0708e-06 | 8.0678e-06 | 8.0648e-06 | 8.0618e-06
45 | 0.00079818 | 0.79287 | 1e-05 | 7.9287e-06 | 7.9705e-06 | 8.0904e-06 | 8.1182e-06 | 8.0847e-06 | 8.0818e-06 | 8.0789e-06 | 8.0761e-06 | 8.0732e-06 | 8.0703e-06 | 8.0675e-06 | 8.0646e-06
46 | 0.00080094 | 0.79286 | 1e-05 | 7.9286e-06 | 7.9721e-06 | 8.0922e-06 | 8.1201e-06 | 8.0867e-06 | 8.0839e-06 | 8.0812e-06 | 8.0784e-06 | 8.0757e-06 | 8.0729e-06 | 8.0702e-06 | 8.0675e-06
47 | 0.00080695 | 0.79285 | 1e-05 | 7.9285e-06 | 7.9738e-06 | 8.094e-06 | 8.1221e-06 | 8.0888e-06 | 8.0861e-06 | 8.0835e-06 | 8.0809e-06 | 8.0783e-06 | 8.0756e-06 | 8.073e-06 | 8.0704e-06
48 | 0.00081636 | 0.79284 | 1e-05 | 7.9284e-06 | 7.9757e-06 | 8.0961e-06 | 8.1242e-06 | 8.091e-06 | 8.0885e-06 | 8.086e-06 | 8.0835e-06 | 8.0809e-06 | 8.0784e-06 | 8.0759e-06 | 8.0734e-06
49 | 0.00082937 | 0.79281 | 1e-05 | 7.9281e-06 | 7.9778e-06 | 8.0983e-06 | 8.1266e-06 | 8.0934e-06 | 8.091e-06 | 8.0886e-06 | 8.0862e-06 | 8.0838e-06 | 8.0814e-06 | 8.079e-06 | 8.0765e-06
50 | 0.00084624 | 0.79278 | 1e-05 | 7.9278e-06 | 7.9801e-06 | 8.1007e-06 | 8.1291e-06 | 8.0961e-06 | 8.0937e-06 | 8.0914e-06 | 8.0891e-06 | 8.0868e-06 | 8.0845e-06 | 8.0821e-06 | 8.0798e-06
51 | 0.0008673 | 0.79274 | 1e-05 | 7.9274e-06 | 7.9826e-06 | 8.1034e-06 | 8.1318e-06 | 8.0989e-06 | 8.0967e-06 | 8.0944e-06 | 8.0922e-06 | 8.09e-06 | 8.0878e-06 | 8.0855e-06 | 8.0833e-06
52 | 0.00089296 | 0.7927 | 1e-05 | 7.927e-06 | 7.9854e-06 | 8.1063e-06 | 8.1349e-06 | 8.102e-06 | 8.0998e-06 | 8.0977e-06 | 8.0956e-06 | 8.0934e-06 | 8.0913e-06 | 8.0891e-06 | 8.087e-06
53 | 0.00092373 | 0.79264 | 1e-05 | 7.9264e-06 | 7.9885e-06 | 8.1095e-06 | 8.1382e-06 | 8.1054e-06 | 8.1033e-06 | 8.1013e-06 | 8.0992e-06 | 8.0971e-06 | 8.0951e-06 | 8.093e-06 | 8.091e-06
54 | 0.00096023 | 0.79257 | 1e-05 | 7.9257e-06 | 7.992e-06 | 8.1131e-06 | 8.1419e-06 | 8.1091e-06 | 8.1072e-06 | 8.1052e-06 | 8.1032e-06 | 8.1012e-06 | 8.0992e-06 | 8.0972e-06 | 8.0952e-06
55 | 0.0010032 | 0.7925 | 1e-05 | 7.925e-06 | 7.9959e-06 | 8.1171e-06 | 8.146e-06 | 8.1133e-06 | 8.1114e-06 | 8.1095e-06 | 8.1076e-06 | 8.1057e-06 | 8.1037e-06 | 8.1018e-06 | 8.0999e-06
56 | 0.0010537 | 0.7924 | 1e-05 | 7.924e-06 | 8.0003e-06 | 8.1217e-06 | 8.1506e-06 | 8.118e-06 | 8.1161e-06 | 8.1143e-06 | 8.1124e-06 | 8.1106e-06 | 8.1087e-06 | 8.1069e-06 | 8.105e-06
57 | 0.0011127 | 0.7923 | 1e-05 | 7.923e-06 | 8.0052e-06 | 8.1268e-06 | 8.1558e-06 | 8.1232e-06 | 8.1214e-06 | 8.1196e-06 | 8.1178e-06 | 8.116e-06 | 8.1143e-06 | 8.1125e-06 | 8.1107e-06
58 | 0.0011818 | 0.79217 | 1e-05 | 7.9217e-06 | 8.0108e-06 | 8.1325e-06 | 8.1616e-06 | 8.1291e-06 | 8.1274e-06 | 8.1256e-06 | 8.1239e-06 | 8.1222e-06 | 8.1205e-06 | 8.1187e-06 | 8.117e-06
59 | 0.0012628 | 0.79202 | 1e-05 | 7.9202e-06 | 8.0173e-06 | 8.1391e-06 | 8.1683e-06 | 8.1358e-06 | 8.1341e-06 | 8.1325e-06 | 8.1308e-06 | 8.1291e-06 | 8.1275e-06 | 8.1258e-06 | 8.1241e-06
60 | 0.0013579 | 0.79185 | 1e-05 | 7.9185e-06 | 8.0247e-06 | 8.1467e-06 | 8.176e-06 | 8.1435e-06 | 8.1419e-06 | 8.1403e-06 | 8.1386e-06 | 8.137e-06 | 8.1354e-06 | 8.1338e-06 | 8.1322e-06
61 | 0.0014701 | 0.79164 | 1e-05 | 7.9164e-06 | 8.0333e-06 | 8.1555e-06 | 8.1848e-06 | 8.1523e-06 | 8.1508e-06 | 8.1492e-06 | 8.1477e-06 | 8.1461e-06 | 8.1445e-06 | 8.143e-06 | 8.1414e-06
62 | 0.001603 | 0.7914 | 1e-05 | 7.914e-06 | 8.0433e-06 | 8.1657e-06 | 8.1951e-06 | 8.1626e-06 | 8.1611e-06 | 8.1596e-06 | 8.1581e-06 | 8.1566e-06 | 8.1551e-06 | 8.1536e-06 | 8.1521e-06
63 | 0.0017616 | 0.79111 | 1e-05 | 7.9111e-06 | 8.055e-06 | 8.1777e-06 | 8.2072e-06 | 8.1747e-06 | 8.1733e-06 | 8.1718e-06 | 8.1703e-06 | 8.1689e-06 | 8.1674e-06 | 8.1659e-06 | 8.1645e-06
64 | 0.0019522 | 0.79077 | 1e-05 | 7.9077e-06 | 8.069e-06 | 8.1919e-06 | 8.2215e-06 | 8.189e-06 | 8.1876e-06 | 8.1862e-06 | 8.1848e-06 | 8.1834e-06 | 8.1819e-06 | 8.1805e-06 | 8.1791e-06
65 | 0.0021833 | 0.79034 | 1e-05 | 7.9034e-06 | 8.0857e-06 | 8.2089e-06 | 8.2387e-06 | 8.2062e-06 | 8.2048e-06 | 8.2034e-06 | 8.202e-06 | 8.2006e-06 | 8.1993e-06 | 8.1979e-06 | 8.1965e-06
66 | 0.0024668 | 0.78983 | 1e-05 | 7.8983e-06 | 8.106e-06 | 8.2296e-06 | 8.2594e-06 | 8.2269e-06 | 8.2255e-06 | 8.2242e-06 | 8.2228e-06 | 8.2215e-06 | 8.2202e-06 | 8.2188e-06 | 8.2175e-06
67 | 0.0028187 | 0.78919 | 1e-05 | 7.8919e-06 | 8.1309e-06 | 8.2549e-06 | 8.2849e-06 | 8.2523e-06 | 8.251e-06 | 8.2496e-06 | 8.2483e-06 | 8.247e-06 | 8.2457e-06 | 8.2444e-06 | 8.2431e-06
68 | 0.0032622 | 0.78838 | 1e-05 | 7.8838e-06 | 8.1619e-06 | 8.2864e-06 | 8.3165e-06 | 8.2838e-06 | 8.2826e-06 | 8.2813e-06 | 8.28e-06 | 8.2787e-06 | 8.2775e-06 | 8.2762e-06 | 8.2749e-06
69 | 0.003831 | 0.78735 | 1e-05 | 7.8735e-06 | 8.2011e-06 | 8.3262e-06 | 8.3565e-06 | 8.3237e-06 | 8.3225e-06 | 8.3213e-06 | 8.32e-06 | 8.3188e-06 | 8.3175e-06 | 8.3163e-06 | 8.3151e-06
70 | 0.0045756 | 0.786 | 1e-05 | 7.86e-06 | 8.2516e-06 | 8.3776e-06 | 8.4081e-06 | 8.3751e-06 | 8.3739e-06 | 8.3727e-06 | 8.3715e-06 | 8.3703e-06 | 8.3691e-06 | 8.3679e-06 | 8.3666e-06
71 | 0.0055751 | 0.7842 | 1e-05 | 7.842e-06 | 8.3181e-06 | 8.4451e-06 | 8.4759e-06 | 8.4427e-06 | 8.4416e-06 | 8.4404e-06 | 8.4392e-06 | 8.438e-06 | 8.4368e-06 | 8.4356e-06 | 8.4344e-06
72 | 0.0069574 | 0.7817 | 1e-05 | 7.817e-06 | 8.4078e-06 | 8.5362e-06 | 8.5674e-06 | 8.5339e-06 | 8.5327e-06 | 8.5315e-06 | 8.5304e-06 | 8.5292e-06 | 8.528e-06 | 8.5269e-06 | 8.5257e-06
73 | 0.0089408 | 0.77814 | 1e-05 | 7.7814e-06 | 8.5322e-06 | 8.6625e-06 | 8.6942e-06 | 8.6602e-06 | 8.659e-06 | 8.6579e-06 | 8.6567e-06 | 8.6556e-06 | 8.6544e-06 | 8.6533e-06 | 8.6521e-06
74 | 0.011921 | 0.77282 | 1e-05 | 7.7282e-06 | 8.7102e-06 | 8.8432e-06 | 8.8756e-06 | 8.8409e-06 | 8.8398e-06 | 8.8386e-06 | 8.8375e-06 | 8.8364e-06 | 8.8352e-06 | 8.8341e-06 | 8.8329e-06
75 | 0.016666 | 0.76442 | 1e-05 | 7.6442e-06 | 8.9736e-06 | 9.1107e-06 | 9.1441e-06 | 9.1084e-06 | 9.1073e-06 | 9.1061e-06 | 9.105e-06 | 9.1038e-06 | 9.1027e-06 | 9.1016e-06 | 9.1004e-06
76 | 0.024805 | 0.75023 | 1e-05 | 7.5023e-06 | 9.3762e-06 | 9.5195e-06 | 9.5544e-06 | 9.5172e-06 | 9.516e-06 | 9.5148e-06 | 9.5137e-06 | 9.5125e-06 | 9.5113e-06 | 9.5102e-06 | 9.509e-06
77 | 0.040144 | 0.7242 | 1e-05 | 7.242e-06 | 1.0001e-05 | 1.0153e-05 | 1.0191e-05 | 1.0151e-05 | 1.015e-05 | 1.0149e-05 | 1.0147e-05 | 1.0146e-05 | 1.0145e-05 | 1.0144e-05 | 1.0143e-05
78 | 0.07214 | 0.67276 | 1e-05 | 6.7276e-06 | 1.0927e-05 | 1.1094e-05 | 1.1134e-05 | 1.1091e-05 | 1.109e-05 | 1.1089e-05 | 1.1087e-05 | 1.1086e-05 | 1.1085e-05 | 1.1083e-05 | 1.1082e-05
79 | 0.1386 | 0.5773 | 1e-05 | 5.773e-06 | 1.2023e-05 | 1.2207e-05 | 1.2252e-05 | 1.2204e-05 | 1.2203e-05 | 1.2202e-05 | 1.22e-05 | 1.2199e-05 | 1.2198e-05 | 1.2196e-05 | 1.2195e-05
80 | 0.20013 | 0.50103 | 1e-05 | 5.0103e-06 | 1.259e-05 | 1.2783e-05 | 1.283e-05 | 1.278e-05 | 1.2778e-05 | 1.2777e-05 | 1.2776e-05 | 1.2774e-05 | 1.2773e-05 | 1.2771e-05 | 1.277e-05
81 | 0.13859 | 0.57731 | 1e-05 | 5.7731e-06 | 1.2023e-05 | 1.2207e-05 | 1.2252e-05 | 1.2205e-05 | 1.2203e-05 | 1.2202e-05 | 1.2201e-05 | 1.2199e-05 | 1.2198e-05 | 1.2197e-05 | 1.2195e-05
82 | 0.072127 | 0.67278 | 1e-05 | 6.7278e-06 | 1.0927e-05 | 1.1094e-05 | 1.1135e-05 | 1.1092e-05 | 1.109e-05 | 1.1089e-05 | 1.1088e-05 | 1.1087e-05 | 1.1086e-05 | 1.1085e-05 | 1.1083e-05
83 | 0.040124 | 0.72423 | 1e-05 | 7.2423e-06 | 1.0001e-05 | 1.0154e-05 | 1.0191e-05 | 1.0152e-05 | 1.0151e-05 | 1.0149e-05 | 1.0148e-05 | 1.0147e-05 | 1.0146e-05 | 1.0145e-05 | 1.0144e-05
84 | 0.024778 | 0.75028 | 1e-05 | 7.5028e-06 | 9.3762e-06 | 9.5197e-06 | 9.5548e-06 | 9.5178e-06 | 9.5168e-06 | 9.5159e-06 | 9.5149e-06 | 9.514e-06 | 9.513e-06 | 9.5121e-06 | 9.5111e-06
85 | 0.016631 | 0.76448 | 1e-05 | 7.6448e-06 | 8.9736e-06 | 9.111e-06 | 9.1446e-06 | 9.1092e-06 | 9.1083e-06 | 9.1074e-06 | 9.1065e-06 | 9.1056e-06 | 9.1048e-06 | 9.1039e-06 | 9.103e-06
86 | 0.01188 | 0.77289 | 1e-05 | 7.7289e-06 | 8.7102e-06 | 8.8435e-06 | 8.8762e-06 | 8.8418e-06 | 8.841e-06 | 8.8401e-06 | 8.8393e-06 | 8.8385e-06 | 8.8376e-06 | 8.8368e-06 | 8.8359e-06
87 | 0.0088927 | 0.77823 | 1e-05 | 7.7823e-06 | 8.5322e-06 | 8.6629e-06 | 8.6949e-06 | 8.6612e-06 | 8.6604e-06 | 8.6596e-06 | 8.6588e-06 | 8.658e-06 | 8.6572e-06 | 8.6564e-06 | 8.6556e-06
88 | 0.0069021 | 0.7818 | 1e-05 | 7.818e-06 | 8.4078e-06 | 8.5366e-06 | 8.5682e-06 | 8.535e-06 | 8.5343e-06 | 8.5335e-06 | 8.5327e-06 | 8.5319e-06 | 8.5311e-06 | 8.5304e-06 | 8.5296e-06
89 | 0.0055126 | 0.78431 | 1e-05 | 7.8431e-06 | 8.3181e-06 | 8.4456e-06 | 8.4768e-06 | 8.444e-06 | 8.4433e-06 | 8.4425e-06 | 8.4418e-06 | 8.441e-06 | 8.4403e-06 | 8.4395e-06 | 8.4388e-06
90 | 0.0045058 | 0.78613 | 1e-05 | 7.8613e-06 | 8.2516e-06 | 8.378e-06 | 8.4091e-06 | 8.3766e-06 | 8.3759e-06 | 8.3751e-06 | 8.3744e-06 | 8.3737e-06 | 8.3729e-06 | 8.3722e-06 | 8.3715e-06
91 | 0.0037536 | 0.78749 | 1e-05 | 7.8749e-06 | 8.2011e-06 | 8.3268e-06 | 8.3576e-06 | 8.3253e-06 | 8.3246e-06 | 8.3239e-06 | 8.3232e-06 | 8.3225e-06 | 8.3218e-06 | 8.3211e-06 | 8.3204e-06
92 | 0.0031771 | 0.78854 | 1e-05 | 7.8854e-06 | 8.1619e-06 | 8.287e-06 | 8.3177e-06 | 8.2856e-06 | 8.2849e-06 | 8.2842e-06 | 8.2835e-06 | 8.2828e-06 | 8.2821e-06 | 8.2814e-06 | 8.2807e-06
93 | 0.0027258 | 0.78936 | 1e-05 | 7.8936e-06 | 8.1309e-06 | 8.2555e-06 | 8.2861e-06 | 8.2542e-06 | 8.2535e-06 | 8.2528e-06 | 8.2521e-06 | 8.2515e-06 | 8.2508e-06 | 8.2501e-06 | 8.2494e-06
94 | 0.0023658 | 0.79001 | 1e-05 | 7.9001e-06 | 8.106e-06 | 8.2302e-06 | 8.2608e-06 | 8.2289e-06 | 8.2283e-06 | 8.2276e-06 | 8.227e-06 | 8.2263e-06 | 8.2256e-06 | 8.225e-06 | 8.2243e-06
95 | 0.0020741 | 0.79054 | 1e-05 | 7.9054e-06 | 8.0857e-06 | 8.2097e-06 | 8.2401e-06 | 8.2084e-06 | 8.2077e-06 | 8.2071e-06 | 8.2065e-06 | 8.2058e-06 | 8.2052e-06 | 8.2045e-06 | 8.2039e-06
96 | 0.0018344 | 0.79098 | 1e-05 | 7.9098e-06 | 8.069e-06 | 8.1927e-06 | 8.2231e-06 | 8.1914e-06 | 8.1908e-06 | 8.1902e-06 | 8.1895e-06 | 8.1889e-06 | 8.1883e-06 | 8.1877e-06 | 8.187e-06
97 | 0.0016351 | 0.79134 | 1e-05 | 7.9134e-06 | 8.055e-06 | 8.1785e-06 | 8.2089e-06 | 8.1773e-06 | 8.1767e-06 | 8.1761e-06 | 8.1755e-06 | 8.1748e-06 | 8.1742e-06 | 8.1736e-06 | 8.173e-06
98 | 0.0014675 | 0.79165 | 1e-05 | 7.9165e-06 | 8.0433e-06 | 8.1666e-06 | 8.1969e-06 | 8.1654e-06 | 8.1648e-06 | 8.1642e-06 | 8.1636e-06 | 8.163e-06 | 8.1624e-06 | 8.1618e-06 | 8.1612e-06
99 | 0.0013252 | 0.79191 | 1e-05 | 7.9191e-06 | 8.0333e-06 | 8.1564e-06 | 8.1868e-06 | 8.1553e-06 | 8.1547e-06 | 8.1541e-06 | 8.1535e-06 | 8.1529e-06 | 8.1523e-06 | 8.1517e-06 | 8.1512e-06
100 | 0.0012033 | 0.79213 | 1e-05 | 7.9213e-06 | 8.0247e-06 | 8.1477e-06 | 8.1781e-06 | 8.1466e-06 | 8.146e-06 | 8.1454e-06 | 8.1449e-06 | 8.1443e-06 | 8.1437e-06 | 8.1431e-06 | 8.1426e-06
|  |  |  | I | 1.6924 | 1.6183 | 1.5304 | 1.437 | 1.3551 | 1.2784 | 1.2067 | 1.1397 | 1.0769 | 1.0182 | 0.96324
|  |  |  | Absorbance | -0.011657 | 0.0077897 | 0.032046 | 0.059393 | 0.084893 | 0.11018 | 0.13525 | 0.16008 | 0.18468 | 0.20903 | 0.23312
|  |  |  | Pred. Conc. | 0.013225 | 0.091338 | 0.18877 | 0.29862 | 0.40104 | 0.50262 | 0.60331 | 0.70307 | 0.80188 | 0.89968 | 0.99645
|  |  |  | error | 1.322% | -0.8662% | -1.123% | -0.1384% | 0.1044% | 0.262% | 0.3308% | 0.3072% | 0.1876% | -0.03185% | -0.3549% | 0.65%
|  |  | slope/intercept | 0.24896 | Concentrations => 0: -0.014949
|  |  | linest, zero intercept | 0.2276 | Concentrations => 0: 0
|  |  |  | Pred. Conc. W/0 | -0.051217 | 0.034226 | 0.1408 | 0.26095 | 0.37299 | 0.4841 | 0.59423 | 0.70336 | 0.81143 | 0.91842 | 1.0243
|  |  |  | error | -5.122% | -6.577% | -5.92% | -3.905% | -2.701% | -1.59% | -0.5765% | 0.336% | 1.143% | 1.842% | 2.427% | 3.192% |  | 1
-10 | 0.006519 | 0.78249 | 1e-05 | 7.8249e-06
-9.9 | 0.0066346 | 0.78229 | 1e-05 | 7.8229e-06
-9.8 | 0.0067534 | 0.78207 | 1e-05 | 7.8207e-06
-9.7 | 0.0068753 | 0.78185 | 1e-05 | 7.8185e-06
-9.6 | 0.0070006 | 0.78163 | 1e-05 | 7.8163e-06
-9.5 | 0.0071293 | 0.7814 | 1e-05 | 7.814e-06
-9.4 | 0.0072617 | 0.78116 | 1e-05 | 7.8116e-06
-9.3 | 0.0073977 | 0.78091 | 1e-05 | 7.8091e-06
-9.2 | 0.0075377 | 0.78066 | 1e-05 | 7.8066e-06
-9.1 | 0.0076816 | 0.7804 | 1e-05 | 7.804e-06
-9 | 0.0078297 | 0.78014 | 1e-05 | 7.8014e-06
-8.9 | 0.0079822 | 0.77986 | 1e-05 | 7.7986e-06
-8.8 | 0.0081391 | 0.77958 | 1e-05 | 7.7958e-06
-8.7 | 0.0083007 | 0.77929 | 1e-05 | 7.7929e-06
-8.6 | 0.0084672 | 0.77899 | 1e-05 | 7.7899e-06
-8.5 | 0.0086387 | 0.77868 | 1e-05 | 7.7868e-06
-8.4 | 0.0088155 | 0.77837 | 1e-05 | 7.7837e-06
-8.3 | 0.0089978 | 0.77804 | 1e-05 | 7.7804e-06
-8.2 | 0.0091858 | 0.7777 | 1e-05 | 7.777e-06
-8.1 | 0.0093798 | 0.77736 | 1e-05 | 7.7736e-06
-8 | 0.0095799 | 0.777 | 1e-05 | 7.77e-06
-7.9 | 0.0097865 | 0.77663 | 1e-05 | 7.7663e-06
-7.8 | 0.0099999 | 0.77625 | 1e-05 | 7.7625e-06
-7.7 | 0.01022 | 0.77585 | 1e-05 | 7.7585e-06
-7.6 | 0.010448 | 0.77545 | 1e-05 | 7.7545e-06
-7.5 | 0.010683 | 0.77503 | 1e-05 | 7.7503e-06
-7.4 | 0.010927 | 0.77459 | 1e-05 | 7.7459e-06
-7.3 | 0.011179 | 0.77414 | 1e-05 | 7.7414e-06
-7.2 | 0.011439 | 0.77368 | 1e-05 | 7.7368e-06
-7.1 | 0.011709 | 0.7732 | 1e-05 | 7.732e-06
-7 | 0.011989 | 0.7727 | 1e-05 | 7.727e-06
-6.9 | 0.012278 | 0.77219 | 1e-05 | 7.7219e-06
-6.8 | 0.012578 | 0.77165 | 1e-05 | 7.7165e-06
-6.7 | 0.012889 | 0.7711 | 1e-05 | 7.711e-06
-6.6 | 0.013212 | 0.77053 | 1e-05 | 7.7053e-06
-6.5 | 0.013547 | 0.76993 | 1e-05 | 7.6993e-06
-6.4 | 0.013895 | 0.76932 | 1e-05 | 7.6932e-06
-6.3 | 0.014256 | 0.76868 | 1e-05 | 7.6868e-06
-6.2 | 0.014632 | 0.76801 | 1e-05 | 7.6801e-06
-6.1 | 0.015022 | 0.76732 | 1e-05 | 7.6732e-06
-6 | 0.015428 | 0.7666 | 1e-05 | 7.666e-06
-5.9 | 0.015851 | 0.76586 | 1e-05 | 7.6586e-06
-5.8 | 0.016291 | 0.76508 | 1e-05 | 7.6508e-06
-5.7 | 0.01675 | 0.76428 | 1e-05 | 7.6428e-06
-5.6 | 0.017228 | 0.76344 | 1e-05 | 7.6344e-06
-5.5 | 0.017726 | 0.76256 | 1e-05 | 7.6256e-06
-5.4 | 0.018246 | 0.76165 | 1e-05 | 7.6165e-06
-5.3 | 0.01879 | 0.76069 | 1e-05 | 7.6069e-06
-5.2 | 0.019357 | 0.7597 | 1e-05 | 7.597e-06
-5.1 | 0.01995 | 0.75866 | 1e-05 | 7.5866e-06
-5 | 0.020571 | 0.75758 | 1e-05 | 7.5758e-06
-4.9 | 0.021221 | 0.75645 | 1e-05 | 7.5645e-06
-4.8 | 0.021901 | 0.75526 | 1e-05 | 7.5527e-06
-4.7 | 0.022614 | 0.75403 | 1e-05 | 7.5404e-06
-4.6 | 0.023362 | 0.75273 | 1.0001e-05 | 7.5278e-06
-4.5 | 0.024147 | 0.75137 | 1.0002e-05 | 7.5149e-06
-4.4 | 0.024971 | 0.74994 | 1.0004e-05 | 7.5024e-06
-4.3 | 0.025837 | 0.74845 | 1.0009e-05 | 7.4915e-06
-4.2 | 0.026749 | 0.74688 | 1.0022e-05 | 7.4851e-06
-4.1 | 0.027708 | 0.74523 | 1.005e-05 | 7.4896e-06
-4 | 0.028719 | 0.7435 | 1.0113e-05 | 7.5187e-06
-3.9 | 0.029784 | 0.74168 | 1.0248e-05 | 7.6007e-06
-3.8 | 0.030909 | 0.73976 | 1.0536e-05 | 7.7938e-06
-3.7 | 0.032096 | 0.73774 | 1.1134e-05 | 8.2138e-06
-3.6 | 0.033352 | 0.73561 | 1.2353e-05 | 9.0867e-06
-3.5 | 0.03468 | 0.73336 | 1.4785e-05 | 1.0843e-05
-3.4 | 0.036086 | 0.73099 | 1.954e-05 | 1.4284e-05
-3.3 | 0.037577 | 0.72849 | 2.8644e-05 | 2.0867e-05
-3.2 | 0.039159 | 0.72584 | 4.5713e-05 | 3.318e-05
-3.1 | 0.040839 | 0.72304 | 7.7055e-05 | 5.5714e-05
-3 | 0.042625 | 0.72007 | 0.00013341 | 9.6065e-05
-2.9 | 0.044525 | 0.71693 | 0.00023263 | 0.00016678
-2.8 | 0.046549 | 0.71359 | 0.00040367 | 0.00028806
-2.7 | 0.048708 | 0.71005 | 0.00069233 | 0.00049159
-2.6 | 0.051013 | 0.7063 | 0.0011692 | 0.00082582
-2.5 | 0.053476 | 0.7023 | 0.0019405 | 0.0013628
-2.4 | 0.05611 | 0.69805 | 0.0031611 | 0.0022066
-2.3 | 0.058932 | 0.69353 | 0.0050518 | 0.0035036
-2.2 | 0.061958 | 0.68872 | 0.0079171 | 0.0054526
-2.1 | 0.065205 | 0.68359 | 0.012165 | 0.008316
-2 | 0.068694 | 0.67812 | 0.018326 | 0.012427
-1.9 | 0.072447 | 0.67228 | 0.027062 | 0.018193
-1.8 | 0.076488 | 0.66606 | 0.039174 | 0.026092
-1.7 | 0.080843 | 0.65941 | 0.055586 | 0.036654
-1.6 | 0.08554 | 0.65232 | 0.077315 | 0.050434
-1.5 | 0.090611 | 0.64475 | 0.10541 | 0.067962
-1.4 | 0.096089 | 0.63667 | 0.14087 | 0.089686
-1.3 | 0.10201 | 0.62804 | 0.18453 | 0.11589
-1.2 | 0.10841 | 0.61885 | 0.23694 | 0.14663
-1.1 | 0.11533 | 0.60907 | 0.29821 | 0.18163
-1 | 0.12282 | 0.59866 | 0.36789 | 0.22024
-0.9 | 0.13091 | 0.58761 | 0.44487 | 0.26141
-0.8 | 0.13965 | 0.57591 | 0.5273 | 0.30368
-0.7 | 0.14906 | 0.56355 | 0.61264 | 0.34525
-0.6 | 0.1592 | 0.55056 | 0.69769 | 0.38412
-0.5 | 0.17007 | 0.53695 | 0.77881 | 0.41818
-0.4 | 0.18169 | 0.52277 | 0.85215 | 0.44548
-0.3 | 0.19404 | 0.50811 | 0.91394 | 0.46438
-0.2 | 0.20709 | 0.49307 | 0.9608 | 0.47374
-0.1 | 0.22077 | 0.47778 | 0.99007 | 0.47304
-1.8791e-14 | 0.23495 | 0.46244 | 1 | 0.46245
0.1 | 0.24946 | 0.44724 | 0.99008 | 0.4428
0.2 | 0.26408 | 0.43243 | 0.96084 | 0.4155
0.3 | 0.2785 | 0.41831 | 0.91401 | 0.38234
0.4 | 0.29236 | 0.40517 | 0.85227 | 0.34532
0.5 | 0.30524 | 0.39333 | 0.779 | 0.30641
0.6 | 0.31671 | 0.38308 | 0.698 | 0.26739
0.7 | 0.32629 | 0.37472 | 0.61314 | 0.22976
0.8 | 0.33359 | 0.36848 | 0.52809 | 0.19459
0.9 | 0.33825 | 0.36454 | 0.44608 | 0.16262
1 | 0.34007 | 0.36302 | 0.36972 | 0.13422
1.1 | 0.33897 | 0.36394 | 0.30091 | 0.10952
1.2 | 0.33502 | 0.36727 | 0.24085 | 0.088458
1.3 | 0.32845 | 0.37287 | 0.19009 | 0.070877
1.4 | 0.3196 | 0.38054 | 0.1486 | 0.056548
1.5 | 0.3089 | 0.39003 | 0.11595 | 0.045224
1.6 | 0.29679 | 0.40106 | 0.091401 | 0.036657
1.7 | 0.28374 | 0.41329 | 0.074038 | 0.030599
1.8 | 0.27017 | 0.42642 | 0.062867 | 0.026807
1.9 | 0.25643 | 0.44012 | 0.056882 | 0.025034
2 | 0.24284 | 0.45411 | 0.055114 | 0.025027
2.1 | 0.22964 | 0.46813 | 0.056651 | 0.02652
2.2 | 0.21699 | 0.48196 | 0.060646 | 0.029229
2.3 | 0.20502 | 0.49543 | 0.066314 | 0.032854
2.4 | 0.19381 | 0.50838 | 0.072929 | 0.037075
2.5 | 0.18341 | 0.52071 | 0.079821 | 0.041563
2.6 | 0.17381 | 0.53234 | 0.086384 | 0.045985
2.7 | 0.16503 | 0.54322 | 0.092085 | 0.050023
2.8 | 0.15701 | 0.55333 | 0.096483 | 0.053387
2.9 | 0.14974 | 0.56268 | 0.099238 | 0.055839
3 | 0.14315 | 0.57128 | 0.10013 | 0.057204
3.1 | 0.1372 | 0.57916 | 0.099082 | 0.057385
3.2 | 0.13181 | 0.58639 | 0.096125 | 0.056367
3.3 | 0.12692 | 0.59304 | 0.091422 | 0.054216
3.4 | 0.12245 | 0.59918 | 0.085234 | 0.05107
3.5 | 0.11831 | 0.60491 | 0.077895 | 0.047119
3.6 | 0.11444 | 0.61033 | 0.06978 | 0.042589
3.7 | 0.11074 | 0.61554 | 0.061274 | 0.037716
3.8 | 0.10716 | 0.62064 | 0.05274 | 0.032732
3.9 | 0.10362 | 0.62572 | 0.044496 | 0.027842
4 | 0.10008 | 0.63084 | 0.036798 | 0.023214
4.1 | 0.096502 | 0.63606 | 0.02983 | 0.018974
4.2 | 0.092877 | 0.64139 | 0.023703 | 0.015203
4.3 | 0.089207 | 0.64683 | 0.018462 | 0.011942
4.4 | 0.085508 | 0.65237 | 0.014096 | 0.0091957
4.5 | 0.081803 | 0.65796 | 0.01055 | 0.0069414
4.6 | 0.078124 | 0.66355 | 0.0077405 | 0.0051362
4.7 | 0.074503 | 0.66911 | 0.0055676 | 0.0037253
4.8 | 0.070968 | 0.67458 | 0.0039264 | 0.0026487
4.9 | 0.067545 | 0.67992 | 0.0027152 | 0.0018461
5 | 0.064253 | 0.68509 | 0.0018416 | 0.0012616
5.1 | 0.061106 | 0.69007 | 0.0012255 | 0.00084569
5.2 | 0.058113 | 0.69484 | 0.00080071 | 0.00055636
5.3 | 0.055278 | 0.69939 | 0.00051418 | 0.00035961
5.4 | 0.052602 | 0.70372 | 0.00032511 | 0.00022879
5.5 | 0.050081 | 0.70781 | 0.00020305 | 0.00014372
5.6 | 0.047711 | 0.71169 | 0.00012592 | 8.9618e-05
5.7 | 0.045486 | 0.71534 | 7.8233e-05 | 5.5963e-05
5.8 | 0.043399 | 0.71879 | 4.9367e-05 | 3.5484e-05
5.9 | 0.041441 | 0.72204 | 3.2263e-05 | 2.3295e-05
6 | 0.039605 | 0.7251 | 2.2341e-05 | 1.6199e-05
6.1 | 0.037883 | 0.72798 | 1.6705e-05 | 1.2161e-05
6.2 | 0.036267 | 0.73069 | 1.3571e-05 | 9.9164e-06
6.3 | 0.03475 | 0.73325 | 1.1864e-05 | 8.6995e-06
6.4 | 0.033324 | 0.73566 | 1.0954e-05 | 8.0584e-06
6.5 | 0.031984 | 0.73793 | 1.0479e-05 | 7.7324e-06
6.6 | 0.030722 | 0.74008 | 1.0235e-05 | 7.5749e-06
6.7 | 0.029534 | 0.74211 | 1.0113e-05 | 7.5052e-06
6.8 | 0.028414 | 0.74402 | 1.0054e-05 | 7.4801e-06
6.9 | 0.027357 | 0.74583 | 1.0025e-05 | 7.4768e-06
7 | 0.026359 | 0.74755 | 1.0011e-05 | 7.4839e-06
7.1 | 0.025415 | 0.74918 | 1.0005e-05 | 7.4955e-06
7.2 | 0.024522 | 0.75072 | 1.0002e-05 | 7.5088e-06
7.3 | 0.023676 | 0.75218 | 1.0001e-05 | 7.5225e-06
7.4 | 0.022873 | 0.75358 | 1e-05 | 7.536e-06
7.5 | 0.022112 | 0.7549 | 1e-05 | 7.5491e-06
7.6 | 0.021389 | 0.75616 | 1e-05 | 7.5616e-06
7.7 | 0.020701 | 0.75735 | 1e-05 | 7.5736e-06
7.8 | 0.020047 | 0.7585 | 1e-05 | 7.585e-06
7.9 | 0.019424 | 0.75958 | 1e-05 | 7.5958e-06
8 | 0.018831 | 0.76062 | 1e-05 | 7.6062e-06
8.1 | 0.018264 | 0.76162 | 1e-05 | 7.6162e-06
8.2 | 0.017724 | 0.76256 | 1e-05 | 7.6256e-06
8.3 | 0.017207 | 0.76347 | 1e-05 | 7.6347e-06
8.4 | 0.016714 | 0.76434 | 1e-05 | 7.6434e-06
8.5 | 0.016242 | 0.76517 | 1e-05 | 7.6517e-06
8.6 | 0.01579 | 0.76597 | 1e-05 | 7.6597e-06
8.7 | 0.015357 | 0.76673 | 1e-05 | 7.6673e-06
8.8 | 0.014942 | 0.76746 | 1e-05 | 7.6746e-06
8.9 | 0.014544 | 0.76817 | 1e-05 | 7.6817e-06
9 | 0.014162 | 0.76884 | 1e-05 | 7.6884e-06
9.1 | 0.013796 | 0.76949 | 1e-05 | 7.6949e-06
9.2 | 0.013443 | 0.77012 | 1e-05 | 7.7012e-06
9.3 | 0.013105 | 0.77072 | 1e-05 | 7.7072e-06
9.4 | 0.012779 | 0.7713 | 1e-05 | 7.713e-06
9.5 | 0.012465 | 0.77185 | 1e-05 | 7.7185e-06
9.6 | 0.012163 | 0.77239 | 1e-05 | 7.7239e-06
9.7 | 0.011873 | 0.77291 | 1e-05 | 7.7291e-06
9.8 | 0.011592 | 0.77341 | 1e-05 | 7.7341e-06
9.9 | 0.011322 | 0.77389 | 1e-05 | 7.7389e-06
10 | 0.011062 | 0.77435 | 1e-05 | 7.7435e-06

## Sheet2
Concentration (atom density) vs measured absorbance
atom density
0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
absorbance