COMBINING FR FILES

if you know the freq resp of mic A re mic B, and the freq resp of mic B

re mic C, you can combine those two .fr files to get the freq resp of mic A

re mic C by using the Combine fr files option under the 'other' menu.  the

order the files are specified is important, so if one.fr goes from A to B

and two.fr goes from B to C and you want to go from A to C, then specify

one.fr as the 'first' file and two.fr as the 'second' file.

this feature is handy when you calibrate something with what you think is

a flat mic, only later to find it's not so flat.  the supposed flat one

can be calibrated with a thought-to-be genuinely flat one and then the

old .fr files can be combined.

this feature can also be used to check to make sure the code handles .fr

files in the Mic_Adjust fields correctly.  for example:  take .cal curves

from the earphone mics as well as the probe tube mics.  also take .fr

files comparing the earphone mics and the probe tube mics.  then use

this feature to combine the earphone mic .fr files with the .fr files

just taken so that a third .cal file can be taken with the earphone mics

but inferred to the probe tube mics.

SANITY CHECKS

the procedures described here are mainly for the programmer.  they're

used in debugging to make sure that everything is self-consistent.

.fr files.  run calib.fr with a B&K on the Reference_Channel, a Knowles

on the other channel, and no probe tube files, except perhaps one for

the B&K.  run calib.fr again as before, but with this new .fr file

specified for the probe tube file that the channel the Knowles is on.

both channels should now look identical and be as flat as the ref mic

ear.cal files.  if not done already, create $XDPHYSDIR/$HOST.k[left,

right].fr files using calib.fr, earphone assemblies and probe tubes.

connect both earphone assembly mics and run calib.cal with check off to

create an ear.cal file.  then run it again with check on.  result

should be flat, except for TL and WN phase curves, which naturally

should be random across frequencies (they'd be equal across channels IF

the sounds were synthesized at the same time, but they're not).

all at once.  a method which checks ear.cal and .fr files

simultaneously is to run calib.cal with check off as before.  then

switch over to the probe tube mics (not forgetting to specify the

appropriate .fr files), turn check on, and run it again.  result should

be flat.

note that all thse checks can also be done w/o a B&K;  things just

won't be flat, but the should be identical.

so, to check everything, do the following:

1. position both earphone assemblies in close proximity facing each

other with both speakers and both knowles mics hooked up

2. use calib.fr to create .fr and $XDPHYSDIR/$HOST.k[left, right].fr

files by treating contralateral knowles mic as the reference ptube.

3. use calib.fr again to check the .fr file, as described above

4. use calib.cal to create an ear.cal file.

5. use calib.cal again to check the ear.cal file, as described above.

6. leaving the chamber closed, switch mic leads to simulate measuring from the

probe tubes.  the left ptube file will now have to be entered in the

right mic's filename slot, and vice versa.

7. use calib.cal once again as described above to check everything all

at once.