this module can calibrate both speakers and microphones.  in

both cases, a speaker, which is not necessarily flat, and either

1 or 2 microphones, one of which must either be naturally flat

(ie a B&K) or have a .fr file which can be used to make it flat,

are used.  the channel the flat mic is on is specified in the

Reference_Channel field of the Prefs window.  it's sensitivity

in mV/Pa is specified in the appropriate So field.  if it's not

naturally flat, then a .fr file is specified in the Mic_Adjust

field as well (in this case, the So of the flat mic used to make

the .fr field should be entered in the So field of the Prefs

table).  unlike the cal module, there is no default for

the Mic_Adjust field.

to calibrate a speaker, a single flat mic is used.  data will be

collected from the other channel as well, but it should be

ignored.

to calibrate a microphone, two mics are used.  the mic to be

calibrated is on the channel the flat mic is not.  it's So field

should be set to 0, and it's Mic_Adjust should be blank.